**Open Source Ventilator to Address Actual/Predicted Ventilator Shortage Worldwide due to COVID19 Pandemic – Distribute Freely**

V1.0 – Sem Lampotang; March 17, 2020, Department of Anesthesiology, Center for Safety, Simulation & Advanced Learning Technologies, University of Florida, Gainesville, Florida, USA

**Engineering Design Specifications**

**Overall design philosophy**

* Open source for use worldwide and contributions from others worldwide
* Adult ventilator (older adults at higher risk)
* A bare bones, safe design is better than nothing and what “nothing” means
* The supplies and materials will be locally available at hardware stores
* The design(s) will be validated; validation will be documented and transparent
* The design(s) will be modular allowing different modules to be mixed and matched depending on local availability
* Dissemination via Internet and GitHub(<https://github.com/CSSALTlab/Open_Source_Ventilator.git>) – web pages have to be in Unicode font to eventually support different fonts and languages

**Assumptions:**

1. FDA will waive clearance for the bare bones design if there is a massive shortage
2. Traditional medical components and supplies used in ventilators will be in short supply
3. Transportation will be impaired/disrupted. Locally available non-medical supplies and equipment readily available at local hardware stores such as sprinkler valves and PVC tubing will be used

**Ventilator Specifications:**

|  |  |  |  |
| --- | --- | --- | --- |
|  | **Range** | **Accuracy** | **Settings** |
| **Volume Controlled Intermittent Mechanical Ventilation** |  |  |  |
| Tidal volume | 500 – 1,500 ml | ± 100 ml | 5 fixed settings: 500, 750, 1,000, 1,250, 1,500 ml |
| Respiratory rate | 10 – 30 bpm | ± 2 bpm | 5 fixed settings: 10, 15, 20, 25, 30 bpm |
| I:E ratio; no inspiratory pause | 1:1 or 1:2 |  | 2 fixed settings |
| PEEP | 0 – 30 cm H2O | ± 1 cm H2O | Adjustable or fixed settings depending on design |
| Air and O2 supplied at 50 ± 5 psig (345 kPa) |  |  |  |
| FiO2 | 0.21 – 1.0 | ± 0.05 | 3 fixed FiO2 settings: 0.21, 0.5, 1.0 |
| Anti-asphyxia valve | Opens at –3 ± 1 cm H2O |  |  |
| Over-Pressure Valve | Opens at 60 cm H2O |  |  |
|  |  |  |  |
| **Optional** |  |  |  |
| In-built blender | 0.5 FiO2 |  |  |
| Oxygen sensor | 0.21 – 1.0 FiO2 | ± 0.05 FiO2 |  |

**Modules:**

Pneumatics

Inspiratory valve

Exhalation valve

PEEP valve

Pressure Regulation

3D print

Body

Accessories

Electronics

Printed circuit board

Microcontroller

Software or Timing circuit

Power Supply

12V DC preferred

User Interface